****

**PROJECT REPORT ON**

**GUN OXIMETER**

SUBMITTED TO THE SAVITRIBAI PHULE PUNE UNIVERSITY, PUNE

IN THE PARTIAL FULFILLMENT OF THIRD YEAR OF ENGINEERING

**IN**

**ELECTRONICS AND TELECOMMUNICATION**

**BY**

|  |  |
| --- | --- |
| **VAISHNAVI V. DESHMUKH** | **PRN No. 72023821J** |
| **PRANAV R. ARJUN** | **PRN No. 72023742E** |
| **MALIK TAUFIQUE** | **PRN NO.** **72162968H** |

**UNDER THE GUIDANCE OF**

**Mrs. Mamta Wanjre**

**ACADEMIC YEAR: 2022-23**

**CERTIFICATE**

This is to certify that Project Report entitled

**“*GUN OXIMETER*”**

Submitted by

***Ms. VAISHNAVI V. DESHMUKH* (*PRN No. 72023821J*)**

***Mr. PRANAV R. ARJUN******(PRN No.72023742E*** ***)***

***Mr. MALIK TAUFIQUE (PRN No.*** ***72162968H*** ***)***

is the record of bonafide work carried out by them in partial fulfillment of the requirement for the award of the Degree of **Bachelor of Engineering (Electronics and Telecommunication),** as prescribed by the Savitribai Phule Pune University in the Academic Year 2022-23.

This project report has not been earlier submitted to any other Institute or University for the award of any degree or diploma*.*

|  |  |  |
| --- | --- | --- |
| **Mrs. Mamta Wanjre** | **Dr. M.P. Sardey** | **Dr. P.B. Mane** |
| Internal Guide | Head of Department | Principal |
| **Department of E&TC Engg.** | **Department of E&TC Engg.** | **AISSMS** |
|  |  | **Institute of Information Technology Pune.** |
|  |  |  |
| Date: | . |  |
|  |  |  |

**Acknowledgement**

It is my great pleasure in expressing sincere and deep gratitude towards my guide **Mrs. Mamta Wanjre,** Electronics & Telecommunication Engineering Department for his valuable guidance and Project coordinator **Mrs. Deepali Yewale** for his constant support throughout this work and help to peruse additional studies in electronics and telecommunication.

We take this opportunity to thank Head of the Department **Dr.M.P. Sardey** and all staff members of department of Electronics &Telecommunication Engineering AISSMS IOIT, Pune, for cooperation provided by them in many ways.

The motivation factor for this work was the inspiration given by our honorable principal **Dr. P.B. Mane.**

Lastly, I am thankful to those who have directly or indirectly supported for our work.

Sign Sign Sign

***Ms. VAISHNAVI V. DESHMUKH Mr. PRANAV ARJUN Mr. MALIK TAUFIQUE***

|  |  |  |
| --- | --- | --- |
| **Sr No.** | **TITLE** | **PAGE NO.** |
| **1 a** | **List of Figures** | **5** |
| **1 b** | **Abstract** | **6** |
| **2** | **Introduction** | **7** |
| **3** | **Literature review** | **8** |
|  | Review of research paper |  |
|  | Review of patents |  |
| **4** | **Aim and Objective** | **10** |
|  | Problem statement |  |
|  | Objectives |  |
|  | Methodology |  |
|  | Specification of the system |  |
| **5** | **Block diagram of the system and its explanation** | **13** |
| **6** | **Hardware design** | **14** |
|  | Construction and circuit diagram |  |
| **7** | **Software design** | **20** |
|  | Modules |  |
|  | Diagrams |  |
|  | Feasibility study |  |
| **8** | **Result and analysis** | **29** |
| **9** | **Challenges and Conclusion** | **32** |
| **10** | **Appendix** | **33** |
|  | References & Bill of material |  |
|  | Datasheets |  |

**INDEX**

|  |  |  |
| --- | --- | --- |
| **FIGURE NO** | TITLE | **PAGE NO** |
| 1 | Gun Oximeter | 6 |
| 2 | Temperature Sensor Principle | 11 |
| 3 | Spo2 Sensor Principle | 12 |
| 4 | Block Diagram of system | 13 |
| 5 | Data Flow diagram | 13 |
| 6 | Arduino UNO R3 | 14 |
| 7 | MLX90614 Sensor | 15 |
| 8 | MAX30102 Sensor | 16 |
| 9 | 20x4 LCD | 16 |
| 10 | Hardware Circuit | 18 |
| 11 | Circuit Diagram | 19 |
| 12 | Django Workflow | 21 |
| 13 | Interfacing Hardware with Software | 27 |
| 14 | Feasibility Study | 28 |
| 15 | Readings using Hardware | 29 |
| 16 | Temperature Accuracy Graph | 30 |
| 17 | Oxygen Accuracy Graph | 30 |
| 18 | MYSQL Database Workbench | 31 |

**LIST OF FIGURES**

**ABSTRACT**


Gun oximeter is a IOT based project which can we very useful in pandemic and post pandemic period. The COVID-19 virus has changed the lifestyle of people in every way. But coping with this virus we have to take precautions.

High temperature and Low oxygen which are the most primary signs needs to checked frequently at public places. For this Gun Oximeter which is a combination of Temperature Gun and Oximeter , will be very useful.

Gun Oximeter contains two main sensors , IR temperature sensor and Spo2 sensor which helps us detect two parameters using single device.

More over we have integrated this hardware with a Web Application. Web Application totally changes the picture of traditional measuring systems. The sensor data which is collected at public places is displayed on a web page . Here one can store the data along with the person’s information.

The sensor data + person’s information is stored in an organized way so that this data can be further used for any type of analysis.



**Fig.1** Gun Oximeter Circuit

**INTRODUCTION**

Deadly corona virus changed our lifestyle in every way. We are stuck in our houses. Once we step out for work we are in constant danger of getting infected. But we have to find a feasible way to live with this new lifestyle. While adjusting to the rush one must take into account that Corona Virus is not yet over.

And neither it will be totally eradicated. Government are making several attempts to make sure every person is aware of the basic things which can help us protect from getting affected.

High Temperature and Low Oxygen level are primary signs of infected person.

Often times people are not aware that they are infected unless their O2 levels are checked.

At the public places only temperature is checked. Oxygen level is totally neglected. Oxygen level is more important as it tells the severity of the infection. There are the cases where due to normal temperature but very low oxygen level, people are getting into serious conditions.

There are several places where currently measuring oxygen saturation level is ignored. This is due to several problems

1. Problem of handling two instruments at same time.
2. Cost of the two devices is more.
3. There is also carelessness among the people.
4. Unawareness among the people.

However even though if the temperature is measured , no proper record is kept. A hard copy is used by the security guards , which is not a proper way of storing the data. There is actually no use of measuring the temperature as the data is neglected.

This data is very useful while making the data analysis of covid patients. One can use this data to check which area is in red zone , which age group is most affected etc.

Gun Oximeter is a device where one can measure both temperature and oxygen at public place . Not only this we can store this sensor data in organized way in a database along with the person’s information which can be retrieved easily.

**LITERATURE REVIEW**

1. **A Novel Patient Monitoring System Using Photoplethysmography and IOT in the**

**Age of COVID-19**

Due to covid outbreak . Thus, there has been a constant demand for a technology that is reliable, inexpensive, simple, and at the same time portable for the clinical settings for community and primary care. The wide availability of small and low-cost semiconductor components, with an advancement in the computer-based pulse wave analysis techniques provide a great platform to achieve this. The performance of this type of device and the criteria for screening are worth studying.

1. **Contactless Temperature Detection of Multiple People and Detection of Possible**

**Corona Virus Affected Persons Using AI Enabled IR Sensor Camera**

The detection of temperature can be made possible by the use of the computer leveraging vision techniques which is equipped with the Raspberry-pi camera module. The process is based on the thermal imaging of the person which can detect the elevated temperature of the person and prevents them from entering into the institution or offices thereby the spread due to the possibly affected persons can be avoided thereby the spread can be controlled. The system not only identifies the person with high temperature but also checks whether the person is wearing a mask or not. The real time analysis of the system is the major advantage of the proposed system.

1. **Current COVID-19 Analysis and Future Pandemics Prediction using Internet of Everything (IoE)**

The oximetric measurement of oxygen concentration in blood has been a valuable tool since it became commercially available in the United States in the early 1970's. Generally, an oximeter is a photoelectric instrument that continually measures the oxygen content of blood or oxygen saturation in a person by measuring the intensity of a light beam transmitted through body tissue. Oxygen saturation is numerically displayed as a percentage, and is typically accompanied by an audible alarm if the current value is outside present limits of acceptable saturation. Motion artifact continues to be a significant source of error and false alarms

1. **Pulse Rate and Blood Oxygen Monitor to Help Detect Covid-19: Implementation and Performance**

Developed and tested the performance of an Atmel ATmega 328P MCU and MAX30100 sensor kit-based pulse rate and blood oxygen monitor. We applied the device on 12 subjects and compared the data with a commercially available, Rossmax SB150 pulse oximeter. A minimal deviation of 0.8175% for pulse rate and 0.425% for blood oxygen was obtained that endorses the accuracy of our algorithm and implementation.

1. **COVID-19 Self-Detection Magic Mirror With IoTbased Heart Rate and Temperature Sensors**

Proposed a solution to let the user do a COVID-19 self-detection by using Magic Mirror with IoT-based technology. This Magic Mirror uses two sensors (heart rate and temperature sensor) to measure user’s heart rate and body temperature. If the user is suspected of having COVID-19, an alert will be displayed on the Magic Mirror or smartphone to let the user take further necessary action

1. US6129673A - Infrared thermometer - Google Patents

To overcome the limitations of the prior art, the invention provides a method and system for performing non-contact temperature measurement of an object accurately and efficiently.

In addition to the ambient T a , the first thermistor 19 is positioned within the chamber to be exposed to the waveguide 15 and, consequently, to IR radiation entering the probe 4 through the window 17.

1. DE19526556A1 - Infrared temperature sensor - Google Patents

Infrared temperature sensors advantageously do not require any Surface contact, but instead use an infrared detector, the infrared light gets from the place where a temperature measurement should be done. Appropriate electronics with the infrared sensor is connected, determines the temperature at the desired location.

1. US4883353A - Pulse oximeter - Google Patents

BACKGROUND OF THE INVENTION This invention relates to displaying the process used by a pulse oximeter for identifying the maximum and minimum values of pulsatile waveforms in order to determine the amplitudes used for calculating oxygen saturation.

Under most operating conditions, the upper envelope and lower envelope will correctly track the peaks and valleys of the pulsatile waveform, indicating an accurate processing of saturation of oxygen percentile.

A pulse oximeter in accordance with claim 2 wherein said displayable process waveform comprises: an upper envelope for tracking the process of determining said maximum value of the pulsatile waveform; and a lower envelope for tracking the process of determining said minimum value of the pulsatile waveform.

1. EP2750604B1 - Wearable pulse oximetry device - Google Patents

TECHNICAL FIELD The present invention relates to systems and methods for pulse oximetry measurements at the wrist, particularly, the present invention relates to a pulse oximetry device that can be worn on a wrist.An increased optical path length, as provided by the configuration of the oximeter 30, brings about an increase of interaction between the propagating light and surrounding tissue for ultimately providing a robust signal from which pulse oximetry data can be obtained.

In other embodiments of the present technique, the wrist-band type oximeter utilizes a system for detecting and/or mitigating signal artifacts arising out user motion, for example, hand motion, thereby achieving a reliable pulse oximetry measurement.

**AIM AND OBJECTIVE**

**Problem Statement:** To design a single device which can measure the temperature and oxygen of people in public places. Storing this data in organized way on database for further analysis.

**Aim**

The motive of Gun Oximeter is to make the process of measuring temperature and oxygen of people at public places much easier , faster and safer. Along with this storing the data collected from public places in such organized way so that it can be used for data analysis in future.

**Objectives:-**


1. To design two in one device for accurate measurement of temperature and oxygen .
2. To interface the various sensors for detection with the microcontroller.
3. To collect the sensor data (temperature, oxygen saturation readings.) in proper format.
4. To build a Web Application where the sensor data will be stored along with person’s information.
5. To migrating this data in Database where it can be stored properly.
6. To interface hardware with software webapp, collecting the sensor readings and storing it for further analysis.

**METHODOLOGY**

1. The project is divided into two parts. Hardware and Software.
2. Gun Oximeter works on two main sensors. IR Temperature detection sensor (MLX90614) and SPO2 Oxygen saturation sensor (MAX30102).
3. Both the sensors are interfaced with the Arduino Board . Measurements of the person ; temperature and oxygen readings are taken on the Arduino IDE.
4. The temperature and Spo2 reading are collected from the sensors and are then sent to the Python terminal.
5. Here the interfacing of Arduino with Python is done. The readings taken from sensors are displayed on the python terminal.
6. In this project Web Application is built using Django Framework. Here the readings are sent to a form.
7. This form contains several other fields related to person’s information such as name, age, area , vaccination status etc.
8. Once the form is filled completely the data of the person along with the readings are saved in a database.
9. This allows us to use the information for further analysis.


**SPECIFICATION OF THE SYSTEM: -**

1. **Temperature sensor (MLX90614):**



**Fig 2** Temperature sensor principle

We have used MLX90614 Infrared thermometer. It has a lens to focus the infrared light. This light is emitting from the object onto a detector called thermopile. Thermopile is a thermocouple , when radiation falls on its surface it gets absorbed and converts into heat. Equivalent voltage output is produced.

1. **Spo2 sensor (MAX30102):**

**Fig 3** Oximeter Principle

Pulse oximetry is simple mechanism. It has two different light sources and a [photodiode](https://www.sciencedirect.com/topics/engineering/photodiode) .Based on the measurement site there are two types of modes. One is transmissive and other is reflective mode. In the transmissive mode, the light sources and photodiode are opposite to each other. The measurement site between them. Light is passed through the site. The reflective mode has different principle. The light sources and photodiode are on the same side, and light is reflected to the photodiode across the measurement site.


**Detailed Specifications:**

1. Operating Voltage= 5V
2. Operating temperature = Room temperature.
3. Spo2 sensor **MAX30102** =
   1. Power supply 3.3V to 5.5V
   2. Current draw ~600μA (during measurements)


~0.7μA (during standby mode)

* 1. IR LED Wavelength 880nm
  2. Red LED Wavelength 660nm

1. Temperature sensor **MLX90614** =
   1. Operating Voltage: 3.6V to 5V (available in 3V and 5V version)
   2. Supply Current: 1.5mA
   3. Object Temperature Range: -70° C to 382.2°C
   4. Ambient Temperature Range: -40° C to 125°C
2. Efficiency = 40% to 60% more than static solar panel.
3. Communication Interface= USB or any external power source
4. Tracking Accuracy= -5% to +5% error can be induced.

****

**Fig 4** BLOCK DIAGRAM OF THE SYSTEM



**Fig 5** Data Flow Diagram

**Hardware Design**

In this project we have used following hardware components:

1. Arduino UNO R3 board.
2. Temperature Sensor (MLX90614)
3. SPO2 Sensor (Max30102)
4. 20x4 LCD Display
5. Breadboard
6. Jumper Wires.

**1.Arduino Uno R3**:-

Arduino UNO R3 microcontroller board is used in this project. ATmega328 AVR microcontroller is used in Arduino UNO .


* Arduino has total 20 input output pins.
* It has two serial pins RX and TX.
*  Programs can be loaded on to it from the easy-to-use Arduino computer program

**Fig 6** Arduino UNO R3

The Arduino has an extensive support community, which makes it a very easy way to get started working with embedded electronics. The R3 is the third, and latest, revision of the Arduino Uno.

* Serial Pins 1 (Tx) and 0 (Rx): Tx and Rx pins are used to receive and transmit TTL serial data.
* Reset Pin: Making this pin LOW, resets the microcontroller.

**2.Temperature Sensor (MLX90614)**:-

The MLX90614 -  **Contactless Infrared (IR) Digital Temperature Sensor .** This can be used to measure the temperature without actual contact of the person. The sensor uses Infrared radiations to measure the temperature of the object. It communicates with I2C protocol.

**Fig 7** MLX90614 Sensor


**MLX90614 Pinout Configuration**

|  |  |  |
| --- | --- | --- |
| **Pin No.** | **Pin Name** | **Description** |
| 1 | Vdd (Power supply) | Vdd can be used to power the sensor, typically using 5V |
| 2 | Ground | The metal can also act as ground |
| 3 | SDA – Serial Data | Serial data pin used for I2C Communication |
| 4 | SCL – Serial Clock | Serial Clock Pin used for I2C Communication |

**MLX90614 Temperature Sensor Specifications**

* Operating Voltage: 3.6V to 5V (available in 3V and 5V version)
* Supply Current: 1.5mA
* Object Temperature Range: -70° C to 382.2°C
* Ambient Temperature Range: -40° C to 125°C
* Accuracy: 0.02°C
* Field of View: 80°
* Distance between object and sensor: 2cm-5cm (approx.)

**3.Oxygen Saturation Detection Sensor (MAX30102)**:-

The MAX30102 is pulse oximetry and heart-rate monitor module. It has internal LEDs, photodetectors, optical elements, and low-noise electronics with ambient light rejection.

It has two LEDs, a photodetector, optimized optics, and low-noise analog signal processing to detect pulse oximetry (SpO2) and heart rate (HR) signals.


Inside the sensor on one side, the MAX30102 has two LEDs – a RED and an IR LED. Other side is a very sensitive photodetector.

**Fig 8** MAX30102 Oxygen Sensor

**Technical Specifications**

|  |  |
| --- | --- |
| Power supply | 3.3V to 5.5V |
| Current draw | ~600μA (during measurements) |
| ~0.7μA (during standby mode) |
| Red LED Wavelength | 660nm |
| IR LED Wavelength | 880nm |
| Temperature Range | -40˚C to +85˚C |
| Temperature Accuracy | ±1˚C |

**4.Liquid Crystal Display (20x4)**:-



A 20x4 LCD display is a very basic module that is commonly used in many devices and circuits. It is preferable to use these modules instead of seven segment and other multi-segment LEDs. This is because LCDs are economical, easily programmable, are capable of displaying custom characters (unlike seven segments), animations, and so on.

A 20x4 LCD can display 20 characters per line and there are 4 such lines. The LCD displays each character as a 5x7 pixel matrix. There are two registers on this LCD, namely Command and Data **Fig** 9 20x4 LCD

The LCD on this controller is standard HD44780.



1. 20 Characters x 4 Lines
2. Built-in HD44780 Equivalent LCD Controller
3. Works directly with ATMEGA, ARDUINO, PIC and many other microcontroller /kits.
4. 4 or 8 bit data I/O interface
5. Low power consumption

**Pin Configuration**


|  |  |  |
| --- | --- | --- |
| **Pin No:** | **Pin Name:** | **Parameters** |
| Pin#1 | It denoted as Vss | It is ground pinout potential at this pinout is zero. |
| Pin#2 | It denoted as Vdd | At this pinout, five volts are provided. |
| Pin#3 | This pinout denoted as Vo | This pinout is used to set the contrast of the screen. |
| Pin#4 | This pin denoted as RS | It used to H/L register select signal. |
| Pin#5 | It denoted as R/W | It used for H/L read/write signal. |
| Pin#6 | This pinout denoted as E | It used for H/L enable signal. |
| Pin#7-14 | The pinouts from seven to fourteen denoted as DB0 – DB7. | It used for H/L data bus for 4 bit or 8-bit mode. |
| Pin#15 | It identified as A (LED+) | It used to set backlight anode. |
| Pin#16 | It recognized as K (LED-). | It used to set backlight cathode. |

**WORKING OF GUN OXIMETER.**

* In the Gun oximeter, Ultrasonic sensor detects the Human Presence. Once a Object is detected at certain distance Gun Oximeter is switched on.
* After switching on , when object is at particular distance, temperature sensor gets activated and temperature is displayed on LCD display.
* Then after temperature , person is asked to get more closer to the gun oximeter for SPo2 detection. After standing at particular distance , SPO2 sensor gets activated and spo2 level is displayed.
* The distances for which the sensors should switched on and off can be decided by user.after the data is being captured it is then sent to the python server.
* The Python server will then send the data to the Django Web page .
* Once the temperature and spo2 readings are displayed on the webpage the user can then enter the person’s information.
* After the form is filled it is submitted . This data is saved in the MySQL database.
* This data can be used for future uses.



**Fig 10** Hardware Circuit

****

**Fig 11** Circuit Diagram

**Software design**

We created the web site using Django Framework also we can display and store the data

What is the Django Framework?

Django is open-source python web framework used for rapid development, maintainable, clean design, and secure websites. A framework is a toolkit of all components needed for application development. We can use this component directly

The goal of the Django framework is to allow developers to focus on the main components of the application . Django is fully featured than many other frameworks on the market. It take cares all other backend components and enables users to focus on developing components needed for their application.

Top 5 Uses of Django

1. Easy to Use

2.It’s fast and simple

3. Excellent Documentation for real-world application


4. It’s secure

5. It suits any web application project

**Django MTV architecture components:**


* Model: it defines the logical data structure.. All the classes representing the logical structure of the database are stored in a script named models.py.
* View: it defines the business logic, in the sense that it communicates with the model and translates it into a format readable by the Template. A view is a Python function, which takes a request as input and returns a Web response as output.
* Template: it defines the structure or the layout of a file, such as a HTML file. It is a text document or a Python string encoded through the Django template language.

The following figure illustrates the Django MTV architecture and how the Model, Template and View components interact each other:




**Fig 12** Django Work Flow

This is our website which we have created . In this web site we have created different fields for persons information and also we have interface hardware readings with software so we displayed the readings on our website for temperature and oxygen . For person information we have created fields like name , age, address etc.



We have also connected data base to our web site . We have used my SQL database for storing the data . When we submit the form then all the data of person with readings of temperature and oxygen will be stored in the database . Thus in this structured manner we can easily store the data .We can use this data for further analysis .



**Code**

**1. Arduino Code :**

//GUN OXIMETER.

// Adding libraries

#include <Wire.h>

#include "MAX30105.h"

#include "spo2\_algorithm.h"

#include <Adafruit\_MLX90614.h>

#include <LiquidCrystal.h>

int Contrast=145;

LiquidCrystal lcd(12, 11, 5, 4, 3, 2);

int temp;

#include<NewPing.h>

const int trigPin = 9;

const int echoPin = 10;

float duration, distance;

Adafruit\_MLX90614 mlx = Adafruit\_MLX90614();

MAX30105 particleSensor;

#define MAX\_BRIGHTNESS 255

#if defined(\_\_AVR\_ATmega328P\_\_) || defined(\_\_AVR\_ATmega168\_\_)

//Arduino Uno doesn't have enough SRAM to store 100 samples of IR led data and red led data in 32-bit format

//To solve this problem, 16-bit MSB of the sampled data will be truncated. Samples become 16-bit data.

uint16\_t irBuffer[50]; //infrared LED sensor data

uint16\_t redBuffer[50]; //red LED sensor data

#else

uint32\_t irBuffer[50]; //infrared LED sensor data

uint32\_t redBuffer[50]; //red LED sensor data

#endif

int32\_t bufferLength; //data length

int32\_t spo2; //SPO2 value

int8\_t validSPO2; //indicator to show if the SPO2 calculation is valid

int32\_t heartRate; //heart rate value

int8\_t validHeartRate; //indicator to show if the heart rate calculation is valid

byte pulseLED = 11; //Must be on PWM pin

byte readLED = 13; //Blinks with each data read

void setup() {

//LCD Set UP

analogWrite(6,Contrast);

lcd.begin(20, 4);

//Presence Sensor Setup

pinMode(trigPin, OUTPUT);

pinMode(echoPin, INPUT);

Serial.begin(9600);

// //HR and SPO2 Sensor Setup

Serial.begin(9600); // initialize serial communication at 115200 bits per second:

pinMode(pulseLED, OUTPUT);

pinMode(readLED, OUTPUT);

if (!particleSensor.begin(Wire, I2C\_SPEED\_FAST)) //Use default I2C port, 400kHz speed

{

Serial.println(F("MAX30105 was not found. Please check wiring/power."));

while (1);

}

Serial.read();

byte ledBrightness = 60; //Options: 0=Off to 255=50mA

byte sampleAverage = 4; //Options: 1, 2, 4, 8, 16, 32

byte ledMode = 2; //Options: 1 = Red only, 2 = Red + IR, 3 = Red + IR + Green

byte sampleRate = 100; //Options: 50, 100, 200, 400, 800, 1000, 1600, 3200

int pulseWidth = 411; //Options: 69, 118, 215, 411

int adcRange = 4096; //Options: 2048, 4096, 8192, 16384

particleSensor.setup(ledBrightness, sampleAverage, ledMode, sampleRate, pulseWidth, adcRange);

//Configure sensor with these settings

mlx.begin();

}

void loop() {

digitalWrite(trigPin, LOW);

delayMicroseconds(1000);

digitalWrite(trigPin, HIGH);

delayMicroseconds(1000);

digitalWrite(trigPin, LOW);

duration = pulseIn(echoPin, HIGH);

distance = (duration\*.0343)/2;

lcd.clear();

lcd.setCursor(0, 0);

lcd.print(distance);

lcd.print(",");

// Temperature

temp= mlx.readObjectTempC();

//LCD

delay(500);

lcd.clear();

lcd.setCursor(0, 2);

lcd.print("temperatre= ");

lcd.print(mlx.readObjectTempC());

lcd.print("C");

delay(1000);

//SPO2

long irValue = particleSensor.getIR();

bufferLength = 50; //buffer length of 100 stores 4 seconds of samples running at 25sps

//read the first 50 samples, and determine the signal range

for (byte i = 0 ; i < bufferLength ; i++)

{

while (particleSensor.available() == false) //do we have new data?

particleSensor.check(); //Check the sensor for new data

redBuffer[i] = particleSensor.getRed();

irBuffer[i] = particleSensor.getIR();

particleSensor.nextSample();

}

maxim\_heart\_rate\_and\_oxygen\_saturation(irBuffer, bufferLength, redBuffer, &spo2, &validSPO2, &heartRate, &validHeartRate);

//Continuously taking samples from MAX30102. Heart rate and SpO2 are calculated every 1 second

while (1)

{

//dumping the first 25 sets of samples in the memory and shift the last 75 sets of samples to the top

for (byte i = 15; i < 50; i++)

{

redBuffer[i - 15] = redBuffer[i];

irBuffer[i - 15] = irBuffer[i];

}

//take 25 sets of samples before calculating the heart rate.

for (byte i = 25; i < 50; i++)

{

while (particleSensor.available() == false) //do we have new data?

particleSensor.check(); //Check the sensor for new data

digitalWrite(readLED, !digitalRead(readLED)); //Blink onboard LED with every data read

redBuffer[i] = particleSensor.getRed();

irBuffer[i] = particleSensor.getIR();

particleSensor.nextSample();

if (validSPO2){

Serial.print(distance);

Serial.print(",");

Serial.print(temp);

Serial.print(",");

Serial.print(spo2, DEC);

Serial.println(",");

}

delay(1000);

//LCD

lcd.setCursor(0, 3);

lcd.print("spo2= ");

lcd.print(spo2);

lcd.print("%");

}

maxim\_heart\_rate\_and\_oxygen\_saturation(irBuffer, bufferLength, redBuffer, &spo2, &validSPO2, &heartRate, &validHeartRate);

}

}

}

**Python Code**

import serial

import time

import schedule

from django.http import JsonResponse

import numpy as np

def Arduino(request):

arduino = serial.Serial('com3', 9600)

print('Established serial connection to Arduino')

data = arduino.readline()

data = str(data,'utf')

data=data.strip('\r\n')

data=data.split(",")

# Decalring the varibales



dist=(data[0])

temp = (data[1])

spo2 = (data[2])


arduino.close()

return JsonResponse({'temp': temp, 'spo2': spo2})



**Fig 13** Web Page and Interfacing hardware with software.

**Feasibility Study:**




**Fig 14** Feasibility Study

**Test Results and Analysis**

We have taken 10 readings using Gun Oximeter. About 90% of the temperature readings were accurate . The accuracy of temperature sensor of our device is quite good . The readings of the Spo2 sensor fluctuated at times. Invalid readings were captured.

Accuracy of temperature sensor is better than Spo2 sensor.

For the software part we used Django as Web framework and My SQL as Database.



**Fig 15** Readings using hardware



**Fig 16** Temperature accuracy graph


**Fig. 17** Oxygen accuracy graph

The 10 readings were saved in the Database as shown below.

In these 10 readings accuracy of temperature sensor is very high as compare to spo2 sensor.

Spo2 sensor readings are fluctuating as you can see reading no 7.

Temperature sensor readings were accurate.



**Fig 18** My SQL Database Workbench

**Challenges and Problems faced**

1. The sensors not showing accurate readings at times due to fluctuations.
2. Delay in outputs. (eg MAX30102 takes 50 samples)
3. Looping the functioning of the three sensors simultaneously.
4. Power and memory management of prototype and Arduino board.

**CONCLUSION**

We have built an IOT based project . This project is divided into two main parts . Hardware and Software. We have built a device which can measure temperature and oxygen . Moreover this data will be store on MySql database. We have Arduino as our microcontroller to interface with the sensors.

For the storing part we have developed a webpage which can store the person’s readings along with the information. This data will be further analyzed for future uses.

**References**

Quast, S.; Kimberger, O. The Significance of Core Temperature—

Pathophysiology and Measurement Methods; Dräger Medical GmbH:

Lübeck, Germany, 2014.

[2] http://www.ces.fau.edu/nasa/module-2/correlation-between-

temperature-and-radiation.php

[3] Hsuan-Yu Chen, Andrew Chen, Chiachung Chen, Investigation of the

Impact of Infrared Sensors on Core Body Temperature Monitoring by

Comparing Measurement Sites, MDPI, Basel, Switzerland, May 2020

Quast, S.; Kimberger, O. The Significance of Core Temperature—

Pathophysiology and Measurement Methods; Dräger Medical GmbH:

Lübeck, Germany, 2014.

[2] http://www.ces.fau.edu/nasa/module-2/correlation-between-

temperature-and-radiation.php

[3] Hsuan-Yu Chen, Andrew Chen, Chiachung Chen, Investigation of the

Impact of Infrared Sensors on Core Body Temperature Monitoring by

Comparing Measurement Sites, MDPI, Basel, Switzerland, May 2020

1. A Novel Patient Monitoring System Using Photoplethysmography and IOT in the Age of COVID-19 Anamika Chauhan Information Technology,Delhi Technological . Kunal Farma Information Technology,Delhi Technological Abhay Goel Information Technology.
2. Contactless Temperature Detection of Multiple People and Detection of Possible Corona Virus Affected Persons Using AI Enabled IR Sensor Camera Abirami M.1, Saundariya K.2, Senthil Kumaran R.3, and Yamuna I.4 1UG Scholar, 2IEEE Member, 3Associate Professor, 4Assistant Professor Department of Electronics and Communication Engineering, IFET College of Engineering, Villupuram, India
3. Current COVID-19 Analysis and Future Pandemics Prediction using Internet of Everything (IoE) Shivangi Sharma Ajay Kumar Singh M.Tech (CSE) Dept. Of CSE Computer Science and Engineering Computer Science and Engineering
4. Pulse Rate and Blood Oxygen Monitor to Help Detect Covid-19: Implementation and Performance Navid Bin Ahmed Dept of EEE Independent Univ, Banglaesh Dhaka, Bangladesh
5. COVID-19 Self-Detection Magic Mirror With IoTbased Heart Rate and Temperature Sensors Maria Seraphina Astriani Computer Science Department, Faculty of Computing and Media Bina Nusantara University Jakarta, Indonesia 11480 seraphina@binus.ac.id Andreas Kurniawan Computer Science Department, Faculty of Computing and Media Bina Nusantara University Jakarta, Indonesia 11480

**APPENDIX**

\

Quast, S.; Kimberger, O. The Significance of Core Temperature—

Pathophysiology and Measurement Methods; Dräger Medical GmbH:

Lübeck, Germany, 2014.

[2] http://www.ces.fau.edu/nasa/module-2/correlation-between-

temperature-and-radiation.php

[3] Hsuan-Yu Chen, Andrew Chen, Chiachung Chen, Investigation of the

Impact of Infrared Sensors on Core Body Temperature Monitoring by

Comparing Measurement Sites, MDPI, Basel, Switzerland, May 2020

**A1 Bill of Material:**

1. Arduino Board-459/-
2. MLX90614 temperature sensor – 1500/-
3. LCD – 200/-
4. MAX30102 SPO2 sensor – 500/-
5. Distance Sensor- 300/-
6. Bread Board – 100/-
7. Jumper wires-100/-

**A2 Datasheets:**

1. **MAX30102 Datasheet**

**[** [**https://datasheets.maximintegrated.com/en/ds/MAX30102.pdf**](https://datasheets.maximintegrated.com/en/ds/MAX30102.pdf)**]**

1. **MLX90614**

**[https://www.sparkfun.com/datasheets/Sensors/Temperature/MLX90614\_rev001.pdf]**

1. **Arduino UNO R3**

**[https://www.farnell.com/datasheets/1682209.pdf]**